| Document Title: Calibration – pH Meter | Document Control No.: [Enter Document Control No.] | Rev. No.: [Enter Revision No.] |
| --- | --- | --- |
| Department: [Enter Department] | Supersedes: Original | Effective Date: [Enter Date] |
| Prepared by: [Affix Name, Signature, Position Title] | Reviewed by: [Affix Name, Signature, Position Title] | Approved by: [Affix Name, Signature, Position Title] |
| Date: [Enter Date] | Date: [Enter Date] | Date: [Enter Date] |

- Purpose:

  - To lay down a standard procedure for calibration of pH meter.

- Frequency:

  - Each day of use, before and after maintenance.

  - [ADD FREQUENCY HERE]

- Tools and Equipment:

  - pH meter

  - pH electrode with temperature probe

  - 50mL beakers

  - 200mL waste beakers

  - Magnetic stir plate and stir bars (optional)

  - Personal Protective Equipment (PPE)

    - Laboratory coat

    - Gloves (e.g., nitrile, latex, vinyl)

    - Safety glasses

    - Gas mask

    - Earplugs

    - Laboratory shoes

  - Camera (Photos for reference)

  - [ENTER LIST OF EQUIPMENT HERE]

- Consumable:

  - Purified water (e.g., deionized [DI], reverse osmosis [RO], or distilled)

  - pH calibration buffers

  - pH electrode cleaning solution

  - pH electrode storage solution

  - [ENTER LIST OF CONSUMABLES HERE]

- Responsibility:

  - Quality Control Head

    - To provide training to the concerned person before the implementation of the SOP.

    - To ensure the procedure is followed.

  - Quality Control Officer/Personnel

    - Ensure responsible for the operation, calibration, and maintenance of the pH meter and its related documentation.

  - Quality Assurance Officer/Personnel

    - To review and check the SOP.

    - To ensure that the procedure is followed.

  - Plant Head and Quality Head

    - To take in-charge of the approval of the SOP.

  - [ADD OR CUSTOMIZE RESPONSIBLE PERSON HERE]

- Procedure/Instructions:

  - Calibration of pH meter

    - Preparation of Reference Buffer Solution

      - Pour about 30 mL of pH 1.68, 5.01/5.00, 7.00/7.01, 10.00/10.01, and 12.45 buffer into a 50 mL beaker and cover the beaker with a watch glass or parafilm before calibration.

      - Pour about 30 mL each of the pH .68, 5.01/5.00, 7.00/7.01, 10.00/10.01, and 12.45 buffers into separate 50 mL beakers. Use these three beakers as the rinse beakers during calibration.

      - Allow all buffers to reach the same temperature since pH readings are temperature-dependent.

    - Calibration

      - Make sure to select the correct measurement before starting the calibration.

      - Wash the electrode thoroughly with purified water or a rinse solution.

      - Do not wipe the electrode as this causes electrostatic charge build-up on the glass surface.

      - Maintain the temperature of the buffers to 25 ± 2 °C, unless otherwise specified in the individual monograph, and dip the pH electrode with the temperature probe into the buffers.

      - Perform the five-point calibration using the prepared reference buffers.

      - Dip the electrode into the calibration buffer.

      - Completely immerse the end of the electrode into the sample.

      - Stir the electrode gently to create a homogeneous sample.

      - Press the key to enter pH calibration mode.

      - Observe the primary display where the measured reading is shown. The smaller secondary display indicates the pH standard buffer solution reading depending on the unit.

      - Wait for the measured pH value to stabilize.

      - Press the key to confirm calibration. The meter is now calibrated to the current buffer.

      - First, rinse the electrode with purified water, followed by the next buffer solution, and place it in the buffer solution.

      - Follow steps 6.1.2.6 to 6.1.2.13 for the additional calibration points.

      - Observe if all the calibration points are placed in the unit configuration setup. This means that the meter returns to the measurement mode automatically.

      - Record the calibration details and temperature on the pH Meter Calibration Record Form.

      - Change the calibration buffer every week or whenever required.

      - Record every measurement required during the scheduled calibration on the pH Meter Calibration Record Form.

  - [ADD OR CUSTOMIZE PROCEDURE HERE]

- Corrective Action:

  - If the slope is out of range (< 95% or > 105%), conduct the following corrective actions in sequential order (if the first corrective action resolves the problem, continuing with subsequent corrective actions is not necessary):

    - Change the standard pH buffers and re-calibrate the instrument.

    - Clean the electrode with the pH Electrode Cleaning Solution (follow manufacturer’s instructions).

    - Record all troubleshooting activities under the notes section of the pH Meter Calibration Record Form.

    - Call the manufacturer if the problem persists or cannot be determined or corrected.

  - [ADD OR CUSTOMIZE CORRECTIVE ACTION HERE]

- Reference Document:

  - Operation Manual supplied by the manufacturer.

  - SOP: Handling out calibration results.

  - SOP: Maintenance of Laboratory Instrument/equipment

  - SOP: Analytical instrument/equipment qualification

  - SOP: Instrument/equipment/equipment usage log book.

  - pH Meter Calibration Record

  - [ADD OR CUSTOMIZE REFERENCE DOCUMENT HERE]

- History/Changes:

| Revision No.: | Revision Date: | Description of Change: | Originator / Author Name: | Title / Department: |
| --- | --- | --- | --- | --- |
| 0 | YYYYMMDD | Original | ENTER AUTHOR NAME | ENTER TITLE/ENTER DEPARTMENT |

*Should you need assistance in customizing this document, please email services@consultareinc.com