| Document Title: Calibration – Thermometer | Document Control No.: [Enter Document Control No.] | Rev. No: [Enter Revision No.] |
| --- | --- | --- |
| Department: [Enter Department] | Supersedes: Original | Effective Date: [Enter Date] |
| Prepared by: [Affix Name, Signature, Position Title] | Reviewed by: [Affix Name, Signature, Position Title] | Approved by: [Affix Name, Signature, Position Title] |
| Date: [Enter Date] | Date: [Enter Date] | Date: [Enter Date] |

- Purpose:

  - To describe the procedure used to point-check the accuracy of the thermometers.

- Frequency:

  - Depending on production volumes and the use of monitoring CCP values or SOP values.

  - Any thermometer that has been dropped or abused and is taken out of service.

  - Any “loose" thermometers, or thermometers that have been out of calibration for less than 3 consecutive months.

  - [ADD FREQUENCY HERE]

- Tools and Equipment:

  - Thermometer (i.e., dial-faced thermometer)

  - NIST traceable test thermometer, certified annually by a contractor (optional).

  - Coffee maker

  - Personal Protective Equipment (PPE)

    - Laboratory coat

    - Gloves (e.g., nitrile, latex, vinyl)

    - Safety glasses

    - Gas mask

    - Earplugs

    - Laboratory shoes

  - Camera (Photos for reference)

  - [ENTER LIST OF EQUIPMENT HERE]

- Consumable:

  - Cooked food product, for sampling

  - Crushed ice or distilled water

  - [ENTER LIST OF CONSUMABLES HERE]

- Responsibility:

  - Plant Head and Quality Head

    - To take in-charge of the approval of the SOP.

  - Quality Assurance Officer/Personnel

    - To review and check the SOP.

    - To ensure that the procedure is followed.

  - Quality Control Head

    - To provide training to the concerned person before the implementation of the SOP.

    - To ensure the procedure is followed.

  - Quality Control Officer/Personnel

    - To ensure the responsibility for the operation, calibration, and maintenance of the thermometers and their related documentation.

  - [ADD OR CUSTOMIZE RESPONSIBLE PERSON HERE]

- Procedure/Instructions:

  - General Protocol for Calibration

    - Assign each thermometer with an identification (ID) number.

    - Thermometers intended for measuring higher temperature items, such as cooked products, are calibrated in hot water, while those used for measuring lower temperatures are calibrated in ice water.

    - All thermometers will be calibrated within ±2°F.

    - Thermometers in use are checked against a certified thermometer during calibration, if available. Otherwise, all thermometers are calibrated either against each other or against a thermometer that is used only during calibration. These methods have required a minimum of three thermometers for accuracy.

    - Dial thermometers are not calibrated on both the high and low ends of the range indicated on the dial.

  - Calibration in Ice Water

    - Add crushed ice and distilled water to a clean container to form a watery slush.

    - Place thermometer probe into slush for at least one minute, taking care to not let the probe contact the container.

    - If the thermometer does not read between 30 and 34°F, adjust to 32° F. Nonadjustable thermometers will be removed from use until they have been professionally serviced. Thermometers that have been adjusted for 3 consecutive months will be replaced.

    - Record the results, using actual values, on the thermometer calibration log, along with the date and initials of the person performing the calibration procedure.

  - Calibration in Hot Water

    - Heat a clean container of water to a temperature range that is used for the cooked product. Running clean water through the coffee maker gives a water temperature of approximately 145°F. Another option is to bring a clean container of water to a rolling boil, generally 210-212°F.

    - Place the thermometer probe into the hot water, along with the certified thermometer and/or reference thermometer, for at least one minute, taking care not to let the probe contact the container.

    - If the test thermometer does not read within ±2°F of the reference thermometer, adjust accordingly. Nonadjustable thermometers will be removed from use until they have been professionally serviced. Thermometers that have been adjusted for 3 consecutive months will be replaced.

    - Record the results, using actual values, on the thermometer calibration log, along with the date and initials of the person performing the calibration procedure.

  - [ADD OR CUSTOMIZE PROCEDURE HERE]

- Corrective Action:

  - Retrain any employee found not following the procedures in this SOP.

  - For an inaccurate, bimetallic, dial-faced thermometer, adjust the temperature by turning the dial while securing the calibration nut (located just under or below the dial) with pliers or a wrench.

  - For an inaccurate, digital thermometer with a reset button, adjust the thermometer according to the manufacturer’s instructions.

  - If an inaccurate thermometer cannot be adjusted on-site, discontinue using it, and follow the manufacturer’s instructions for having the thermometer calibrated.

  - Retrain employees who are using or calibrating food thermometers improperly.

  - [ADD OR CUSTOMIZE CORRECTIVE ACTION HERE]

- Reference Document:

  - Thermometer Standardization Record

  - Thermometer Calibration Log

  - Point-Check Calibration of Thermometers Record Form

  - Certificate of Verification (COV)

  - Certificate of Guarantee (COG)

  - Certificate of Conformance (COC)

  - Calibration Certificate

  - [ADD OR CUSTOMIZE REFERENCE DOCUMENT HERE]

- History/Changes:

| Revision No.: | Revision Date: | Description of Change: | Originator / Author Name: | Title / Department: |
| --- | --- | --- | --- | --- |
| 0 | YYYYMMDD | Original | ENTER AUTHOR NAME | ENTER TITLE/ENTER DEPARTMENT |

*Should you need assistance in customizing this document, please email services@consultareinc.com